Implanted cardiac device deactivation and decontamination

We have had 2 incidents in the past 2 years when contaminated Pacemakers and ICDs have been returned to the Clinical Investigations Department without deactivation or decontamination.

- In September 2012 involved an ICD that had not been deactivated and was sent back to the Sterile Services Department of Yeovil hospital through the post.  The member of staff in the sterile services department did not realise that the device had not been deactivated and came into contact with the wires and received a shock.  Following this incident a new protocol for the safe removal and disposal of such devices was put in place and sent to all local funeral directors and community hospitals.

- In August a pacemaker was explanted by a GP at a local funeral directors.  The above procedure was not followed and the contaminated pacemaker was sent via the internal post, wrapped in a latex glove to the Clinical Investigations Department.  Fortunately the staff member opening the post realised what the device was and immediately sent it for decontamination.

We would like to remind all GPs who explant devices post mortem to follow the above procedure at all times for the safety of hospital staff.